## Slide 1: Some Components of QA/QC in Soil Testing at NCDA&CS
David H. Hardy
MASTPAWG 2023
February 22, 2023

## Slide 2: Trained Staff
Standard Operating Procedures – all lab aspects
  New staff shadow and train with senior chemist
Ethics training and sign-off
Safety training and job safety analysis at each station

## Slide 3: Possible Mistakes / Errors
Samples out of order
Transposing lab numbers
Wrong samples pulled off shelves
Worn scoops
Wrong scoops used
Double scoops
Filter paper failure
ICP carryover
ICP missed vials
Standards / extractants problems
ICP tubing plugs
Contamination of blank / stds
Faulty pH electrode
Pumps not calibrated
Contaminated lab water
Cold lab temperatures
Wrong client gets report

## Slide 4: Check Soil- 1 / 36 samples

## Slide 5: Sample Order / Handling

## Slide 6: Correct Samples and Correct Scoops
260 sets per series – 3 series

## Slide 7: Sample Order & Labeling
Mehlich 3
Humic Matter
VW / pH - BpH

## Slide 8: Color & Labeling to Minimize Mistakes / Confusion

## Slide 9: Traceability and Accountability

## Slide 10: Traceability and Accountability

## Slide 11: ICP Setup – Speed Component
Blank- 5% nitric acid
Standards in 5% nitric acid
  Low P, S (Low Std.)
  K, Ca, Mg, Na, Mn, Cu, Zn, Al, Fe (Std 1)
  P & S (High Std2)
Low QC (ICV, low QC)
QC (initial) – ICV
Continuous QC (CCV)
Check soil random in set

## Slide 12: Carefully Set ICP Element Limits

## Slide 13
Orange – out of range
Warning limit
Red – out of limit
Error limit
Green – Data in LIMS
Check sample

## Slide 14: pH / BpH Stds- 4, 7, 10

## Slide 15: Check Sample Preparation

## Slide 16: Check Sample

## Slide 17: Daily Control Charts

## Slide 18: Validation of Data - Agronomist Rechecks
Research samples – at least 10% reruns
Flags identify concern – soil test data relationships

## Slide 20: Validation of Data Agronomist Rechecks
Flags identify the concern
  Soil test data limits
Client reports with soil test limit flags
  Held at analyzed status for agronomist review

## Slide 21: Agronomist View of Flags

## Slide 22: Validation of Data Agronomist Rechecks
Flags identify the concern – 3rd day in lab
Client reports with soil test limit flags
  Held at analyzed status for agronomist review
Agronomist selects rechecks - back to lab in batch
  Used in place of grower duplicate

## Slide 23: Selecting Rechecks for Tests

## Slide 24: Validation of Data Agronomist Rechecks
Flags identify the concern as set through soil test data limits
Client reports with soil test limit flags are held at analyzed status for agronomist review
Agronomist selects rechecks and sends back to lab
Agronomist reviews recheck data
  Resolves
  Request further rechecks or runs
Publish all reports for that day

## Slide 25
Recheck Results- ICP Batch
Report # Lab # Element Original Recheck %RD

## Slide 26
Recheck Results- pH Batch
Lab # pH/ BpH Original Recheck %RD Final

## Slide 28: Proficiency Testing Programs
DENR Certified Lab – pH, P, Zn and Cu

## Slide 29: Closing Comments
No lab is perfect……
Processes that safeguard errors
  SOPs
  Training
  Consistency……
Continuously monitoring processes
Investigate and try to remedy concerns
  Ask for help if needed

## Slide 32: Items Discussed
Potential errors
Sample order
Labeling, color, ….
Logs and traceability
Standardization of instruments
ICP limits for extreme nutrient levels
Agronomist daily review of data
Control charts
Proficiency testing & NC Lab Certification

## Slide 34: Traceability Soil Data Viewer